Wellcome-MRC 
Cambridge Stem Cell Institute

Preparing Centrifuges for relocation: Decontamination

| Reason | Title/Department | Signature |  | Date |
| --- | --- | --- | --- | --- |
| Subject Matter Expert: |  |  |  |  |
| Approved By: |  |  |  |  |

Table of Contents

- Purpose                                                                                                3

- Scope                                                                                                    3

- Responsibilities                                                                                  3

- Procedures                                                                                          4

- Related Documents/Forms                                                              5

- Revision History                                                                                  5

- Glossary                                                                                                6

- Appendices                                                                                          6

- Purpose

With the upcoming move to the Jeffrey Cheah Biomedical Centre, formerly known as “Project Capella”, the purpose of this document is to describe the procedures for the decontamination of benchtop centrifuges, making them ready to be handled for relocation. 
With appropriate action taken by the centrifuge users, some of the risks associated with handling the centrifuges can be reduced, increasing the likelihood the removal company may handle these safely.

- Scope

This SOP will inform and instruct Laboratory-based CSCI personnel on what actions to take to correctly decontaminate centrifuges, in order to declare them safe for handling and ready for relocation.

This SOP is solely applicable to CSCI Research Groups based in the Gleeson Building.
This SOP is applicable to centrifuges used for work up to and including Containment Level 2 (CL2).

- Responsibilities

| Resource | Responsibility |
| --- | --- |
| Centrifuge User | Read and understand this SOP. Be familiar with the properties and hazards of the materials involved. Use suitable alternatives if any branded laboratory disinfectants named within this SOP are not used/available. Wear appropriate PPE when handling or working with laboratory disinfectants. Accurately complete a decontamination report form and securely attach to the centrifuge. |

- Procedures

  - Removing and decontaminating interior parts

- Collect Centrifuge stickers from the Laboratory Management team.

- Open the centrifuge lid.

- Remove any buckets and the centrifuge rotor.

- Spray the buckets and rotor and thoroughly clean using 1% Chemgene and then using 70% ethanol.

- Label all of the individual components using the stickers provided. 
This is to ensure all interior components can be matched to their original centrifuge when unpacked at the JCBC.
If more stickers are required, please inform the laboratory management team, located in the facilities office.

- Attach a decontamination certificate to the rotor.

- Leave the centrifuge rotor near to the centrifuge.

- The rotors are to be packed by Harrow-Green, separately from the centrifuge. This is to prevent internal damage to the centrifuge during the move.

- Clean, labelled accessories and buckets should be wrapped in bubble wrap and packed in a box labelled with the centrifuge number.

  - Decontaminating the Centrifuge

- Wipe down the interior drum of the centrifuge, using 1% Chemgene and then using 70% ethanol.

- Switch on the power supply.

- Close the centrifuge lid, so it mechanically locks and secure shut with a strip of autoclave tape.

- Disconnect the centrifuge from the power supply.

- Wipe down the exterior of the centrifuge, using 1% Chemgene and then using 70% ethanol.

- Wipe down the power cable, using 1% Chemgene and then using 70% ethanol.

- Coil the power cable and secure with cable ties.

- Complete a separate decontamination certificate and tape it to the front of the centrifuge.

- Related Documents/Forms

| Document/Form # | Document/Form Title |
| --- | --- |
| CSCI-RA02 | Risk Assessment – Use of Centrifuges |
| - | University of Cambridge Decontamination Form |

- Revision History

| Version No. | Type | Reason | Effective Date |
| --- | --- | --- | --- |
| 1.00 | Initial Release | To outline correct procedures in preparing centrifuges for relocation | March 2019 |
|  | Annual review |  | March 2020 |

- Glossary

  - Abbreviations

SOP – Standard Operation Procedure
CSCI – Cambridge Stem Cell Institute
CL2 – Containment Level 2
JCBC – Jeffrey Cheah Biomedical Centre
PPE – Personal Protective Equipment

- Appendices

None.